## Slide 1
AUTOMATED IMAGING · BONE HEALTH
Measuring Spinal Bone Health
from Routine Hip / Spine X-rays
An automatic, explainable pipeline that finds the L4 vertebra and measures its bone density and texture.
Presented by: Edward Choi · Data Analyst
Audience: Dr. Brent Lanting / Orthopaedic / Hip Surgery
June 30, 2026
Notes: 'I built an automatic tool that measures bone quality of the L4 spine vertebra directly from the X-rays you already take.
No extra scan, no extra dose.
I'll show what it does, how accurate it is, and what we still need to fix before clinical use.'

## Slide 2
THE GOAL
Why measure L4 bone from an X-ray?
Bone quality of the lumbar spine matters for surgical planning and follow-up.
AP hip / spine X-rays are already taken in routine care.
Goal: a fast, repeatable bone-quality number for the L4 vertebra — fully automatic.
L4 is the standard lumbar reference level for bone assessment (as in DXA).
IMPORTANT CAVEAT
This is a relative, X-ray-based proxy of bone density — not a calibrated DXA result (g/cm²).
It is for method demonstration. True BMD needs DXA / QCT or an in-image calibration phantom.
WHAT WE MEASURE ON L4
1) Bone density proxy (image brightness)
2) Bone texture (trabecular pattern, GLCM)
L4 Spinal Bone Measurement · Data Analytics
2
Notes: ‘This is not replacing DXA. I’m extracting a consistent, automatic bone signal from images you already have.
Two things: density (how bright/dense the bone looks) and texture (the internal bone pattern).
But it is a proxy, not calibrated g/cm-squared.'

## Slide 3
OVERVIEW
The pipeline in four steps
1
Input X-ray
Routine AP spine / hip radiograph
2
Find L4
AI outlines vertebrae L1-L5 and selects L4
3
Measure
Density proxy + bone texture on L4
4
Report + Explain
Numbers + heatmaps showing the 'why'
Built on a single AI detector (YOLOv8) trained on 2,000 public spine X-rays; the measurement and explanation code is independent of the detector.
L4 Spinal Bone Measurement · Data Analytics
3
Notes: 'The pipeline in four steps:
take the X-ray,
the AI finds and outlines L4,
we measure bone on that region, and we report the numbers with a heatmap that shows where the AI looked.
Everything after the detector is standard, transparent image math.'

## Slide 4
STEP 1
Finding the L4 vertebra
AI segmentation (YOLOv8-seg)
Trained on 2,000 labeled spine X-rays (public BUU-LSPINE set).
It outlines each vertebra L1-L5 as a separate object.
L4 is selected by its label, with a 4th-from-top backup rule.
Output: a clean L4 outline at full image resolution.
Blue = all vertebrae · Green = selected L4
L4 Spinal Bone Measurement · Data Analytics
4
Notes: 'The AI does instance segmentation - it separates each vertebra.
and then pick L4.
Green is the L4 it chose.
This is the core building block;
the bone measurement sits on top.'

## Slide 5
VALIDATION
The detector is accurate (in-domain)
0.97
Mask mAP@50
0.95
Precision
0.96
Recall
0.76
mAP@50-95
On held-out validation X-rays, the model finds and outlines vertebrae very reliably.
mAP@50 = 0.97 means near-complete agreement with the human labels at a standard overlap threshold.
Right: model predictions on a validation batch.
L4 Spinal Bone Measurement · Data Analytics
5
Notes: 'On images from the same source it trained on, accuracy is high - 0.97 mAP (mean Average Precision). The maximum value is 1
So the detector itself is strong.
The next slide is the honest part: what happens on our real clinical X-rays, which look different.'

## Slide 6
STEP 2
Measuring bone on L4
Two bone read-outs
Density proxy: average brightness inside the L4 region. Denser bone tends to look brighter.
Texture (GLCM): the trabecular bone pattern — contrast, energy, homogeneity.
We use the inner 'cancellous' core of L4 (red ROI) where bone change shows first.
Green = L4 · Red = measurement ROI
L4 Spinal Bone Measurement · Data Analytics
6
Notes: 'From the L4 outline I measure two things:
density (brightness) and texture (the internal bone pattern).
I focus on the inner core, where osteoporotic change appears earliest.
density here is a relative index, not calibrated BMD.'

## Slide 7
REALITY CHECK
Honest test on 42 real clinical X-rays
0.76
Shape quality (Dice)
When it labels L4, the outline is good.
0.28
As-labeled (Dice)
Sometimes it names the wrong vertebra.
THE ONE ISSUE: 'OFF-BY-ONE'
The AI draws a correct vertebra shape, but on these different-looking clinical images it sometimes labels the neighbour as 'L4'.
Right shape, wrong name — on 28 of 42 images.
Why report two numbers?
Separating 'shape quality' from 'naming' tells us exactly what to fix. The segmentation is good (0.76). The gap to 0.28 is purely a labeling problem we can target directly.
L4 Spinal Bone Measurement · Data Analytics
7
Notes: MOST IMPORTANT SLIDE.
I am being transparent. Two numbers on purpose.
The outline quality is good - 0.76. But the as-labeled score is 0.28 because of an anatomical demarcation issue during ground-truth labeling.
Due to my own difficulty in precisely differentiating the L4 boundaries on real clinical X-rays, the vertebra above or below L4 may have been incorrectly designated as L4.
So Right shape, wrong name.
The value of splitting the metric: it proves the segmentation works and isolates a single, fixable labeling issue

## Slide 8
DIAGNOSIS
Seeing the 'off-by-one' clearly
How to read this
Yellow = true L4 (Edward)
Red = what the AI named L4
Green = best-matching vertebra
Top row (worst): red sits one vertebra above the yellow — the off-by-one shift.
Bottom row (best): red, green and yellow line up.
Cause: I am having difficulty naming the L4 bone. I think it's a labeling issue.
L4 Spinal Bone Measurement · Data Analytics
8
Notes: 'This shows the problem visually.
Yellow is the true L4 , I marked it.
Red is the AI's L4. In the worst cases the red is one level too high.
The shapes are right - it is a counting / field-of-view problem, not a segmentation failure.’
I am having difficulty naming the L4 bone.
It is often difficult to distinguish bones in X-ray images.

## Slide 9
EXPLAINABILITY
Can we trust where the AI looks?
Grad-CAM (gradient-based)
A heatmap of the pixels the model used to decide 'this is L4'.
Red / warm = most important.
The attention lands on the L4 body and its internal bone structure — not on background, implants, or soft tissue.
Why it matters: a clinician can verify the AI is reading the right anatomy before trusting the number.
L4 Spinal Bone Measurement · Data Analytics
9
Notes: 'Trust matters in medicine.
This heatmap shows the pixels the AI used.
It focuses on the L4 bone itself and its trabecular pattern - the same thing a radiologist would look at.
So when it does pick the right level, it is looking at the right structure for the right reason.'

## Slide 10
WHERE WE STAND
What works, and what's next
WORKING TODAY
Accurate vertebra detector (mAP@50 = 0.97 in-domain).
Good L4 outline quality on clinical images (Dice = 0.76).
Fully automatic density + texture read-out.
Explainable: heatmaps confirm correct anatomy.
Honest, two-number evaluation.
NEXT STEPS
1) Fix L4 labeling: fine-tune on a small set of labeled clinical X-rays (removes off-by-one).
2) Calibrate density to real units using a phantom or paired DXA.
3) Validate against DXA on the same patients.
4) Package as a one-click report for routine use.
L4 Spinal Bone Measurement · Data Analytics
10
Notes: 'The engine works and is explainable.
One fixable gap remains - the L4 labeling on clinical images.
The fix is straightforward: fine-tune on a small number of labeled clinical X-rays.
I would need a small set of labeled clinical cases from your team to do step one.'
However, I plan to receive the dataset and label it first. I will ask for your help later.
And I would like to compare this with the DXA test results. It would be helpful to know how actual bone density can be measured.
For now, 2 and 3 numbers are outside the scope of the research, so they have been marked this way.

## Slide 11
Takeaways
We can measure L4 bone density and texture automatically from X-rays you already take.
The detector is accurate and explainable; the L4 outline quality on clinical images is good (0.76). ( 0.97 on the training dataset )
One clear, fixable gap remains: correctly labeling L4 on out-of-domain clinical images.
With a small labeled clinical set + calibration, this can become a routine bone-quality readout.
Thank you — questions & discussion welcome.
Edward · Data Analyst
Notes: 'In short:
I can get an automatic, explainable bone-quality signal from routine X-rays.
The method is sound;
one labeling gap is left and it is fixable with a small amount of labeled clinical data.
My ask: help me get a few labeled clinical cases so I can close that gap.

## Slide 12
Sample Application Description version
L4 Spinal Bone Measurement · Data Analytics
10
Patients Area
Patient X-ray upload history
Graph of bone density changes according to X-ray upload date
Original X-ray image and extracted vertebrae and L4 marking
Extracted L4 vertebra
X-ray Metadata
Bone tissue value based on brightness (0 - 1)
Explainable
AI values
Areas
affecting
BMD
Areas
affecting
L4
Extraction

## Slide 13
Sample Application
L4 Spinal Bone Measurement · Data Analytics
10